## Slide 1
Couch & Couch motor
1D platform
RealSense depth camera
a)
b)

## Slide 2: Couch Tracking Flowchart –with Intel Realsense depth camera UDPsender_realtime.py
Couch
Input Trace
Couch Motor
Depth camera
Frame No. Timestamp, ROI1, ROI2
Platform moves in 1D, based on the input trace.
Depth camera measures depth values in real-time. Write to csv file, frame number, timestamp, ROI1 depth values (1D platform), ROI2 depth values (couch)
Couch is receiving ROI1 depth values via UDP in real-time. Motor travel in opposite direction of 1D platform to achieve motion compensation.
*** For code details please see couch tracking pseudocode document.
1D platform

## Slide 3: Couch Tracking with pre-recorded trace Flowchart – UDPsend_trace.py
Input Trace
Depth camera
Frame No. Timestamp, ROI1, ROI2
UDP sender
Couch Motor
Depth camera
Frame No. Timestamp, ROI1, ROI2
1D platform
Couch

## Slide 4
Intel RealSense depth camera
(Surface tracking)
Couch Compensation
KIM
(Internal motion tracking)
Other methodologies…

## Slide 5
Tracking Target
(6Dof Robotic arm/
1D motion platform)
Movable Couch
Tracking Tool
(KIM/
Depth measurement)
Target being tracked,
Positional measurement generated
Measurements are sent to couch via UDP
Isocenter
Target motion range
Compensated motion range

## Slide 6
KIM
(Generated measurements)
Router
Ethernet connection
Ethernet connection
Integrated Raspberry Pi & DC motor drive
Monitor
Keyboard & Mouse
USB Connection
Motor power supply
Linear actuator motor
Couch
6 degree-of-freedom
robotic arm
Phantom
Elekta/Varian kV imager
kV images in real-time
Track internal motion
Clinical experiment hardware setup
Router establishes private UDP communication with pre-set fixed IP address
Robotic arm moves phantom to simulate patient internal motion
Phantom positions are tracked by kV imager, measurements are generated by KIM algorithm
Pi receives measurements and operate couch to perform motion compensation

## Slide 8
Router
Ethernet connection
Ethernet connection
Integrated Raspberry Pi & DC motor drive
Monitor
Keyboard & Mouse
USB Connection
Motor power supply
Linear actuator motor
Couch
6 degree-of-freedom
robotic arm
Phantom
Intel RealSense Depth Measurement
Track surface motion
Laboratory experiment hardware setup
Router establishes private UDP communication with pre-set fixed IP address.
Robotic arm moves phantom to simulate patient internal motion.
Phantom surface positions are tracked by depth camera.
Pi receives measurements and operate couch to perform motion compensation.

## Slide 9
Robotic arm to replicate target motion
Couch
Linear actuator motor
Intel Realsense depth camera

## Slide 10
Raspberry Pi connection
Original design but being updated
Notes: Add a picture of Pi in Blacktown

## Slide 11
5V
Motor power supply
-
Motor
+
Raspberry Pi

## Slide 13
Ground
5V
Motor
Motor power supply
+
-
1D motion platform

## Slide 14
Target Tracking
Couch
Input motion trace
Couch Feedback
Updated positions
KIM noise (optional)
KIM latency
UDP receive
UDP send
Compensation algorithm
Couch move
Couch positions
Motion plot in real-time
Couch error (optional)
KIM + Couch emulator
Couch latency
50-70ms

## Slide 15
Target Frame Position (input motion trace)
While timestamp is within input motion duration:
Couch operation (latency added)
Get couch position
If couch did not move:
couch position = 0
xxxx
Target isocenter position = Target frame position + couch position
KIM latency
Send delayed target isocenter position to couch algorithm
if couch did move:
Target isocenter position = Target frame position + couch position
KIM latency
Send delayed target isocenter position to couch algorithm
Log and print (frame number, frame timestamp, target isocenter position, Target frame position, couch position)
timestamp+=1
Get current time (time.time())
if current time – start time < Time period x num time periods:
time.sleep(1)
get current time
Target isocenter position (input motion trace + couch position)
Couch and Target position are updated sequentially
Simple to implement, less flexibility

## Slide 16: Multi-threading
More realistic simulation as couch and KIM system are independent in real-world
Couch operation should not affect KIM measurement being generated
More flexibility adding more components on different threads
Clean logic

## Slide 17
Real target position (ground truth)
Timestamp, target position (robot/motion platform input trace, 0.2s interval)
DONE
2. KIM measurement (latency occurs)
Real target position_timestamp, local timestamp after delay, KIM measurement (interpolated from real target position)
3. Couch running log
[Local timestamp when motor starts moving, received KIM measurement,
error (KIM measurement – isocentre), couch position it should travel to)]
Logic from Paul and Ann are fundamentally the same.
Why multi-threading is better

## Slide 18
Frame No. Timestamp, ROI1, ROI2
Frame No. Timestamp, ROI1, ROI2
Frame No. Timestamp, ROI1, ROI2
Frame No. Timestamp, ROI1, ROI2
Frame No. Timestamp, ROI1, ROI2
Frame No. Timestamp, ROI1, ROI2
Frame No. Timestamp, ROI1, ROI2
Frame No. Timestamp, ROI1, ROI2
Target tracking
Couch response
Couch moving
Couch moving
Couch moving
Couch Compensation – Always compensate on latest measurement